## Slide 1: SG9000 OS Setting
SG9000II.ini
SPJT file
AutoLoad Setting

## Slide 2: SG9000.ini
The ini file put in the root directory: C:\SG9000II\
Sg9000.ini file record the os program setting values:
[Tester]
Name=TSG24 //Tester Name
OSVer=V1.16.0 //OS Version
OSVerControl=V1.16.0.0001 //Version Number
Interface=PCI7230 //Set Handler Interface Board(PCI7230,DAQ6509)
SetupPW=OFF //Switch Passwarod function at setup Page
HWScan=OFF //Switch Hardware scan function

## Slide 3: SG9000.ini
[Last] //Record the information by OS about the last project loaded
DeviceName=TestProject2
Customer=Test
TestPlan=TestProject_V1
[Automation] //Set Location of Barcode file for OS to AutoLoad
BarCode = D:\AutoMation\XML\Tester_Barcode.XML

## Slide 4: SPJT File
SPJT File could save OS setting for the project
[DUT] //Record the location of required file for the project
Transform=C:\SG9000II\SYSTEM\Murata\MDFE2PFA029\MDFE2PFA029_revG15_1.trf
Limit=C:\SG9000II\DUTS\Murata\MDFE2PFA029\MDFE2PFA029_revG15_1.lim
Table=C:\SG9000II\DUTS\Murata\MDFE2PFA029\MDFE2PFA029_revG15_1.tab
DLL=C:\SG9000II\DUTS\Murata\MDFE2PFA029\MDFE2PFA029_revG15_1.dll
PinMap =C:\SG9000II\DUTS\MURATA\MDFE2PFA029\MDFE2PFA029_revG15_1.mvp
TPPROFILE = L:\DUTS\Murata\TPPROFILE\MD029
// Transform : transform file (*.trf)
// Limit : Limit Table (*.lim)
// Table : Test Item Table (*.tab)
// DLL : Compiled Test Program
// PinMap : Pin-mapping file (*.mvp) (only in V1.15.1)
// TPPROFILE : Golden File for auto transform

## Slide 5: SPJT File (V1.17.X)
[Setting] //Save setting state at Setup Page for the project
MaxSiteNum=2 //Maxsimun Site Numbers
ProductType=MURATA //ProductType for Lot Name Function
Handler=FT //Handler Type: FT,CP,Manual
ProgramMode=Production //Program Mode: Production,Engineer
SiteFlag=3 //Testing Sites Switch
CSV=ON //CSV format datalog
TXTSummary=OFF //Txt Format datalog
StopOnFail=OFF //Stop On Fail Function
TestTime=OFF //Test Time Function
AutoTransForm=OFF //Auto TransForm Function
SQC=OFF //SQC Function
Backup=OFF //Backup Function

## Slide 6: SPJT File
[CheckStack] //Set Device Stack Check Function
Enabled = 1 //1:On 0:OFF
Stackea = 100 //Set how many device tested to check if stack occur
TestName = Icc0_2450 //Set the test item which have to be monitor
Limit = 0.001 //Set the limit range

## Slide 7: SPJT File (PAT for FT)
[PAT] //Option of PAT for FT
Enable=ON //Turn OFF or ON for PAT function
BinName= PATFail //Set software binname for PAT fail duts
SWBin= 9999 //Set software binnumber for PAT fail duts
HWBin= //Set hardware binnumber for PAT fail duts
PATLimitFile=M: //Indicate where to save PATlimitsFile on Server

## Slide 8: SPJT File (PAT for CP)
[PAT] //Option of PAT for CP
Enable=ON //Turn OFF or ON for PAT function
BinName=PATFailBin //Set software binname for PAT fail duts
SWBin=9999 //Set software binnumber for PAT fail duts
HWBin= //Set hardware binnumber for PAT fail duts
PATLimitFile=P:\TestData\Temp\TSG\CP2 //Indicate where to save PATlimitsFile on Server
CP=YES //Tell OS the PAT is run under CP/FT
RTBin=2,3,4,5,6 //If under wafer sort, Need to set which BIN would be retest.

## Slide 9: SPJT File (GS Rule)
[GSruleSetting]
Enable=on
Time_Hour=48
Warn_Disable=off

## Slide 10: Auto Load Setting
To active Auto Load Funciton:
1. Set Server connect batch file (AutoLoad.bat, AutoRelease.bat)
2. Set Barcode path in SG9000.ini
3. Put setup file at server (Setup filename must be same as set in release form)
4. Put zipped test program at server which location as setup file set.
Example File:
Barcode File
Setup File
AutoLoad.bat
AutoRelease.bat

## Slide 11: Auto Load – Setup File
L:\DUTS\Murata\MD029_FPGA_14G\
MDFE2PFA029_revG17_1.spjt
\DUTS\Murata\MDFE2PFA029\MDFE2PFA029_revG17_1.spjt
1. The Location of zipped test program.
2. Specify the filename of project file to load.
3. Relative directory on local tester of project file.

## Slide 12: Auto Transform (Changeover file)
To activate auto transform function, there are two files required; changeover.txt, golden file(TP profile).
1. changeover.txt
the content of file must have the title as the example at least. As follows:
GU_TestBoard: 130404-2 130404-1
GU_CalDate: 2013-05-10 11:01:01
GU_CalTech: 943408
SD_HandlerID: IS34
GU_CalTray: MD042001010
GU_Status: Site#2 Pass (8 pass, 4 req)

## Slide 13: Auto Transform (Golden File)
2.Golden File:
File name rule: Devicename_GU_bench_xxxxx
The file provide limit range of transform item and raw data of calset.
Example:
-- *********************************************************************
-- **************** Golden Unit LIMIT INFORMATION ************************
-- *********************************************************************
//Limit Range for the test item to add offset
-- *********************************************************************
-- **************** Golden Unit Data **************************************
-- *********************************************************************
//Raw data of golden (calset) that target for added offset values.

## Slide 14: EndLot Script
1. Put a script file name as “Endlot.txt” in the directory of SG9000: \SYSTEM\Device\
OS will execute the exe file that record in the script file after “EndLot”
2. Example File:
3. The files that want to executed in the endlot.txt file have to be put in the same directory as Endlot.txt file.

## Slide 15: Directories
Backup //Store the files that set in backup function
Datalog //Production datalog put in here
DUTs //Test Program and project files
INCLUDE //Library used by OS program
LIB //Library used by Test Program
LSR //LSR files of production
SYSTEM //Some setting and related files for the project
TMP //Temporary files of LSR files
UTILITY //Put some external application for OS